## Slide 1: Software Development and Hardware Commissioning for Applications in the High Contrast Imaging Testbed Facility
Eric Lin
Mentors: Hong Tang and Camilo Mejia Prada

## Slide 2: Introduction
Who?
  Myself, with help from Hong Tang, Camilo Mejia Prada, Keith D. Patterson, Pin Chen, and many more!
What?
  Work on coronagraph testbeds, their instruments, and the vacuum chambers around them
When?
  May 14th – July 19th, 10 week program
Where?
  High Contrast Imaging Testbed Facility (HCIT)
Why?
  To learn new skills and gain some real-life experience
  To help further our ability to study planets beyond our solar system

## Slide 3: Background
What are coronagraphs?
Instrument used to block out direct starlight
Originally developed to study the corona of the Sun
Why do we need them?
Can be used to observe/study exoplanets
Stars are bright, planets are dim
Life beyond our solar system?
What are coronagraph testbeds?
“Million dollar LEGO sets”
Testing new coronagraph technologies and techniques

## Slide 4: Main Project Tasks:
Task 1: Designing and installing a turbo vacuum pump safety system
Task 2: Installing the Decadal Survey Coronagraph Testbed (DST)
Task 3: Tuning the DST’s thermal heater PID controllers
Task 4: Calibrating a Deformable Mirror (DM) mount
Task 5: Mounting, Calibrating, and installing a DM

## Slide 5: Task 1: Turbo Vacuum Pump Outlet Control (TVPOC)
Controls a smart power strip/outlet that a turbo pump and its gate valve will be connected to
  TVPOC polls the pressure transducer 1 time per second, also pump speed/temp if the new pump comes in soon
  If TVPOC detects an error while polling (e.g. pressure too high), TVPOC remotely turns off the outlet the gate valve is connected to
  Turbo pump is designed to automatically shut the gate valve when power is lost, sealing the chamber. Only works when there is enough air pressure in the pump lines.
Smart power strip for TVPOC

## Slide 6: Task 1: TVPOC (Cont.)
User calls tvpoc.py
Date/time of start logged into tvpoc_log.txt
Pressure is polled from pressure transducer
Pressure Too High
Command sent to shut off turbo pump gate outlet
Pressure Normal
Program sleeps for *1 second
LOOP
START
Keyboard Interrupt
Date/time of stop logged
Turbo pump loses power, gate valve closes
* Can be adjusted based on how quickly an error needs to be caught
END

## Slide 7: Task 1: TVPOC (Cont.)
Come up with design of system alongside Camilo
Find and order necessary hardware
Figure out how to communicate with new hardware
Install new hardware
Test and debug various issues
Present system and receive feedback
Alter system design and apply changes
Implement new system into existing Watchdog
Turbo Gate Valve of HCIT1
New setup of the turbo pump
Backing pump originally located
next to the turbo pump

## Slide 8: Task 2: Installing DST
Moving DST into chamber
D-Sub cabling all the motors
Taping up wires
Installing heaters and thermal sensors
Untangling, coiling, recoiling, optical fibers
Balancing the table floaters
Installing coolant lines
Testing everything

## Slide 9: Task 3: Tuning DST thermal PID controllers
PID Control = Calculate error from setpoint given parameters
Working with MATLAB GUI
Estimated P values through manual testing
  Found to be extremely high, much higher than expected
Calculated I and D using P
Values led to uncontrolled behavior
Bug found in C script while debugging another issue
Fixed!

## Slide 10: Task 3: Tuning DST thermal PID controllers (Cont.)

## Slide 11: Task 4: Calibrating a DM mount
Characterizing voltage controlled crystal piston motors
  Controls tilt of mirror (very fine)
Used interferometer to measure tilt
Used dummy mirror so I don’t get blacklisted
  Originally intended to use polished steel plate
  Switched to aluminum mirror
Cycled voltages through motors
Left to do:
  Calibrate strain gauges to eliminated hysteresis (lag)
  Determine modal frequencies (vibrations) of mount

## Slide 12: Task 4: Calibrating a DM mount (Cont.)

## Slide 13: Task 5.1: Mounting a DM
A lot more frustrating and tedious than it should have been
  Strain reliever too thick
  Good tape type ran out
  Copper strips for thermal conductivity
  Working with JCI
Lots of on the fly problem solving and improvising
Learned a lot for next time

## Slide 14: Task 5.2: DM Installation
Preparation of all components
Connecting cables
Working on cooling system

## Slide 15: Task 5 Pictures (Cont.)

## Slide 16: Acknowledgements
Hong Tang
Camilo Mejia Prada
Keith D. Patterson
Pin Chen
All of the HCIT team

## Slide 17: Thank You For Listening!